과 업 지 시 서

사 업 명	2020년 자기공명영상촬영장치(1.5T PHLIPS Achieva) 유지보수 계약
주관기관	충청남도서산의료원

2019.  12.   .

MRI 장비 유지보수 과업지시서

1. 입찰명 : MRI 연간 유지보수

2. 계약기간 : 계약일로부터 ～ 2020.12.31.

3. 유지보수대상 장비 내역

모델명	설치장소	수량	구입일자	장비취득가격
Achieva 1.5T	병원 MRI실	1대	‘08.7.3.	1,150,000,000 (제세비용 제외)

4. 계약조건

가. 이 계약은 위의 장비가 항상 정상상태로 가동될 수 있도록 장비의정기점검 보수작업을 하기 위함에 있다.

나. MRI 본체등 MRI장비와 관련된 모든 구성품 무상공급(COLD HEAD 교체, 정도관리, 조정실 컴퓨터, 항온항습기, 칠러, 헬륨 포함(단, 마그넷과 악세사리(패드)는 제외)

다. MRI 장비 수리와 관련된 교체되는 부품은 정품(제조사의 certification라벨이 붙어 있는 부품을 납품해야 한다.

라. 유지보수 점검자의 자격 : 상기 대상 MRI 제조사의 정비보수 기술 교육을 수료한 certification을 소유한 자(국내거주자로 유지보수 용역을 제공하는 회사의 재직증명 필)

마. 년간 6회(격월) 예방 정기점검으로 진료에 차질이 발생되지 않도록 하고 장애발생시 긴급보수를 실시하여 장비 가동률을 90% 이상으로 유지하도록 하여야 하며 정기점검 실적을 서면으로 통보한 후 유지보수료를 청구하여야 함

바. 년간 6회(격월) 정기 예방점검을 실시하지 않았을 경우에는 유지보수료를 지급하지 않음

사. 점검내역

ㅇ 장비 외관상 이상 유무 확인

ㅇ 안전장치 정상 동작 확인

ㅇ 장비 내부 청소

ㅇ 이미지 이상 유무 확인

ㅇ 마모 부분의 자연소모 보수 및 부품교환 및 조정

ㅇ 기능상 이상 유무 확인

ㅇ 정도관리 시행

아. 본 계약서에 명시되지 않은 사항은 양측이 합의하여 해결한다.

《별첨 1》

MRI 정기점검 내역
No	항 목	비 고
1	주기적 영상점검(PIQT)	▶ 영상점검
2	환자 테이블 점검	▶ 테이블 점검 및 조정
3	중심선 정렬	▶ Laser marker 조정
4	컴퓨터 점검(호스트와 리컨)	▶ 점검및 클리닝
5	코일점검	▶ 외부 및 영상점검
6	주 전원점검	▶ 주 전원 확인
7	노이즈 시험	▶ Tune up
8	인공물의 유,무 시험	▶ Tune up
9	Gradient 점검	▶ Gradient 조정
10	Refrigerator점검	▶ LCC 수량, 온도 점검 및 조정
		▶ Helium Compressror압력점검
11	RF Amp점검	▶ Amp점검 및 조정
12	소프트웨어 점검	▶ 점검및 백업
13	Magnet 점검	▶ 소모품 확인 및 필요시 교체
14	Console 점검	▶ 사용자-환자간 인터콤 확인
		▶ 에러 발생 유무 확인
15	System 점검	▶ 안전장치 이상 유무 확인
		▶ Cooling 장치 확인

《별첨 2》

MRI 유지보수점검보고서

Section		Result	Remarks
1	System level
1.1	Check Phantom	Correct Yes/No
	PIQT Periodic Image Quality Test	Correct Yes/No
1.2	Check stop function	Correct Yes/No
	Check Table movements	Correct Yes/No
	Check helium over-pressure	Correct Yes/No
	Check Gradient airflow sensor	Correct Yes/No
	Gradient coil temperature sensor	Correct Yes/No
	Gradient coil heat exchanger interlocks	Correct Yes/No
	Gradinet Amplifier heat exchanger interlocks	Correct Yes/No
	Visual inspection	Correct Yes/No	No damages etc.
1.3	Fan(s) of the reconstructor	Correct Yes/No
	Dust filter and fans BDAS/CDAS	Exchanged Yes/No Correct Yes/No
	Dust filters and fans MR console	Exchanged Yes/No Correct Yes/No
	Dust filter and fan RF amplifier	Exchanged Yes/No Correct Yes/No
	Dust filter and fan Gradient system	Correct Yes/No
1.4	RF cage	Good/Bad
	Check helium venting system	Correct Yes/No
1.5	Check the RF coils	Correct Yes/No
1.6	Check site conditions	Correct Yes/No	Temperature, Humidity, Dust
1.7	Check for dust in hybrid box and QBC	Correct Yes/No
	Check impedance 50 Ohm QBC load	Correct Yes/No
1.8	Artifact measurements	Correct Yes/No
Section		Result	Remarks
2	Mains distribution
	Check Main power supply voltage	V

Section		Result	Remarks
3	Gradinet System
	Check Gradient coil connections	Correct Yes/No
	Gradient switch box	Correct Yes/No
Section		Result	Remarks
4	Heat Exchanger
	Verifiy the heat exchanger settings	Done Yes/No
	Check water level of Grad. coil heat exchanger	Done Yes/No	N.A.
	Check water level of Grad. amp heat exchanger	Done Yes/No	N.A.
	Check water level of LCC	Done Yes/No	N.A.
Section		Result	Remarks
5	Refrigerator system
	Check pressure of the compressor	Correct Yes/No
	Check water and oil circuit	CorrectYes/No Leak Yes/No
	Check water flow	Correct Yes/No
	Absorber replacement	Done Yes/No	N.A.
	Coldhead maintenance	Done Yes/No	N.A.
Section		Result	Remarks
6	Magnet system
	Check of the helium level and boil off	day/%	< 65cc/h
	Check DPS	Correct Yes/No
	Check the cables at the magnet	Correct Yes/No
Section		Result	Remarks
7	Patient support
	Check the finger protection switch-plate	Correct Yes/No
	Check the patient alarm	Correct Yes/No
	Check the manual switch	Correct Yes/No
	Check the manual decoupling buttons	Correct Yes/No
	Check all moving cables for degradation	Correct Yes/No
	Check patient comfort lights	Correct Yes/No
	Check the emergency stop circuit	Correct Yes/No
	Check patient support cooling fans	Correct Yes/No

Section		Result	Remarks
8	System measurements and adjustments
8.1	Automatic BDAS/CDAS tests	Correct Yes/No
	Automatic BDAS/CDAS tests and calibration	Correct Yes/No
8.2	RF amplifier self test (S23, S24 and S26 only)	Correct Yes/No
	Max kw adjustment and PMU test	Correct Yes/No
	Resonane frequency determination	Correct Yes/No
	Power Reference measurement	Correct Yes/No
	Pickup-coil measurements	Correct Yes/No
8.3	Spike check	Correct Yes/No
Section		Result	Remarks
9	Operators console + viewing
9.1	Monitor adjustment for OCD	Correct Yes/No
9.2	Mechanical maintenance	Correct Yes/No
	Ceiling carriage limit stops	Correct Yes/No
	Functional check	Correct Yes/No
Section		Result	Remarks
10	Computer system
	Cleaning the mouse	Done Yes/No	N.A.
	Internal computer cleaning(when applicable)	Done Yes/No
	Internal computer MOD cleaning(when applicable)	Done Yes/No	N.A.
Section		Result	Remarks
11	Software
	Clean up disks	Done Yes/No
	Backup procedure	Done Yes/No
	Check statistical logging	Correct Yes/No